## Slide 1
ION Pumps for UHV Systems, Cyclotrons Synchrotrons & Particle Accelerators
Mauro Audi , Agilent Vacuum Product
Cyclotrons 2022

## Slide 2
ION Pumps Agilent Technologies
1957-59
Varian Associates invents the first ION Pump

## Slide 3
Closed pump
Do not need any baking pump
No contamination from the roughing line
No moving parts
There is no rotation or moving parts
No lubricant
There are no oil or solvents and is fully contamination free
Can withstand air inrush or improper use
Maintenance free, High reliability
ION Pumps Main choice for UHV systems, cyclotrons , synchrotron & particle accelerators
Why?

## Slide 4
ION Pumps
Basic Pumping Mechanism
Pumping principle:
Plasma discharge in crossed electric and magnetic field act as an electron trap
Ionizing collision between electron and gas
Ion bombardment of titanium cathode
Some ions diffuse into the cathode ( pumped )
Sputtering of chemically active Ti film on anode
Neutral Gas molecules stick to Ti film (chemisorption) and are buried in the anode ( main pumping )
Ion pumps do not pump only ions at the cathode , but mainly neutral molecules at the anode

## Slide 5
ION Pumps
Basic Pumping Mechanism – Noble gases
Ion pumps do not pump only ions at the cathode,but mainly neutral molecules at the anode
This works only for all active (getterable) gases
Noble gases do not react with Ti film
No chemisorption of neutral molecules
Alternative pumping method needed
Some ion bombarding the cathode are neutralized and reflected
They maintain enough energy to be physically implanted into the anode
They will then be covered by the sputtered Ti film on the anode
It is a physical burying, not a chemical reaction
For this reason , pumping much less efficient than for getterable gases

## Slide 6
DIODE ION PUMP
Very limited Noble Gases pumping , both speed and capacity
Low percentage of Ions are reflected and implanted into the anode
Noble gases are only temporarily pumped at cathode
Limited Argon capacity means :
ION Pumps
 | Pure Argon | Air ( 1% Ar)
10 E-5 mbar |  | 8 days
10 E-6 mbar | 20 hours | 3 months
10 E-7 mbar | 8 days | 2.5 years
10 E-8 mbar | 3 months | 25 years
10 E-9 mbar | 2.5 years | 

## Slide 7: Noble Gases Instability
7

## Slide 8
Improve the number/probability of Noble Gas ions reflection vs implantation after bombarding the cathode
After being reflected, they will be physically buried into the anode (no chemical interaction)
Two different approaches to modify Diode to obtain this result:
Both solutions do improve the pumping speed and stability for noble gases
ION Pumps designe for Noble Gases pumping
Different types of pump: Diode, Noble Diode and StarCell
Noble Diode
Replace one Ti Cathode with Ta Cathode
Change of MATERIAL
A
Heavier material increases reflection probability
Triode / StarCell
Design cathodes to obtain a different angle of incidence
Change of GEOMETRY
B
Higher reflection when ions bombard the cathodes at small angles compared to normal incidence

## Slide 9
Noble Diode vs Diode
Much lower capacity and speed for H2 ( Ta vs Ti )
Lower speed for all getterable gases such as N2, CO, CO2 ( Ta vs Ti)
Improved stability for Noble Gases (one order of magnitude)
Argon speed up to 15% 20% of Air Pumping speed vs 5% for a Diode
Starcell vs Diode
Lower speed for all getterable gases (N2, CO, CO2)
Comparable speed for Hydrogen ( Ti vs Ti )
Improved stability for Noble Gases (two orders of magnitude)
Argon speed up to 35% 40% of Air Pumping speed vs 5%
When only getterable gases must be pumped , Diode is still the best choice
ION Pumps
Noble Diode & StarCell vs Diode

## Slide 10
Much higher capacity for H2 ( one order of magnitude )
Similar speed for all getterable gases (N2 , CO , CO2)
Higher stability for Noble Gases ( one order of magnitude)
Higher speed for Noble Gases (almost double , up to 40 % vs 20% of Air )
When Noble gases must be pumped , StarCell is the proper choice
ION Pumps
StarCell vs Noble Diode
Magnet

## Slide 11
In order to reach UHV , outgassing / cleanliness is even more important than pumping speed
The ion pump element must be the cleanest surface exposed to vacuum , because its outgassing is due to both thermal and bombardment induced effect
The cathode surface will eventually be removed (sputtered): therefore not only the surface , but the bulk too must be hydrogen free
Ion pump element is “vacuum fired”
2 hours @ 750 °C , at 1*E-5 mbar
Then, the complete pump is processed under vacuum at 450°C
ION Pumps
Manufacturing

## Slide 12
Pumps are individually processed
Residual gas analysis for each pump
Individual records of each pump ( outgassing , spectra…)
Pumps are processed in nitrogen atmosphere to prevent external oxidation (no more beadblasting needed)
Pumps can be leak-checked at high temperature (Helium instead of Nitrogen)
Vacuum performance are much more repeatable
ION Pumps
Manufacturing

## Slide 13: RGA before process
ION Pumps
Manufacturing - RGA
RGA after process
Thermal Process
13

## Slide 14
Thank you for your attention !