## Slide 1: Measurement of detector response functions for fast neutron spectroscopy with organic scintillators
Tanya Hutton
Andy Buffler, Zina Ndabeni, Kutullo Maibane, Erin Jarvie
CYC2022 Beijing 5-9 December 2022

## Slide 2: Neutron production from cosmic rays
electrons
and photons
pions, protons
and neutrons
muons
cosmic ray
Altitude (km)
Cosmic rays:
  Galactic cosmic radiation (GCR)
  Solar energetic particles (SEP)
Neutrons are produced from evaporation of highly energetic nuclei, knock-on in peripheral collisions, or in charge exchange reactions.
Measured
[Goldhagen et al. Rad. Prot. Dos. (2004) vol.110 p.387]
Calculated
[Clem et al. Rad. Prot. Dos. (2004) vol.110 p.423]
Cosmic-ray neutron spectrum: 20 km at 54° N, 117° W

## Slide 3
In biological systems:
Effects of high-energy neutron radiation
Very little information is available on the effects of high energy neutrons due to a lack of suitable measurement devices.
Concerns for aviation, space missions, workplace exposure, radiation therapy etc.
In electronic systems:
10.21926/rpm.2002012
increased cancer risk, fetal harm etc.
chemical and biological damage
single event upsets, displacement damage, total ionizing dose
soft bit-level errors, hard errors of latch-up and burn-out etc.

## Slide 4: South African Space Neutron Initiative (SASNI)
Geographical importance
Space weather and radiation protection policy (SANSA)
Fast neutron reference beams (iThemba LABS)
Instrumentation for fast neutron spectrometry
Biological effects of fast neutrons

## Slide 5: Organic scintillators
C, H based scintillation medium e.g. EJ-301/BC501A/NE213
Indirect detection via 1H(n, n) and 12C (n, X)
Excellent energy resolution
Fast time response
Sensitive to all radiation types, and capable of discrimination
Typical measurements systems are limited to use in controlled laboratory environment
2” x 2” BC501A
Light output L (MeVee)
Pulse shape S

## Slide 6
Compact (“hand held”)
Flexible analyses
Neutron-gamma separation
Safe to use in extreme environments
Neutron spectrometer over
wide energy range (1-100 MeV)
Flexible suite of applications
Silicon photomultipliers
Digital signal processing
Pulse shape discrimination
Plastic organic scintillators
Measured and simulated mono-energetic response functions
Minerals exploration, smart farming, personal and workplace dosimetry, security, dosimetry in aircraft and spacecraft, …
Concept for new neutron spectrometer

## Slide 7
Detector design
Plastic scintillator
Silicon photomultiplier (SiPM)
Digital data acquisition
EJ-276
MicroFC-60035
DT5730
Pulse shape discrimination
Light output: 56% anthracene
0.6 x 0.6 x 12 cm3
Area: 6 x 6 mm2
18980 microcells
Operating voltage: +28.5 V
14 bit, 0.5 GS/s
0.5-2.0 Vpp
QtDAQ acquisition software
https://github.com/veggiesaurus/qtdaq

## Slide 8: Neutron spectroscopy: time-of-flight
Time of arrival used to determine neutron energy.
Requires ns-pulsed neutron beam, and a sufficiently long flight path (d).
Time of flight channel
Counts
Rad Meas 45 (2010) 1342

## Slide 9
… to produce the neutron energy spectrum
Measured neutron pulse height spectrum …
… is unfolded using a
known detector
response matrix …
Energy (MeV)
Fluence rate (cm-2s-1)
Count rate (s-1)
L (MeVee)
L (MeVee)
Scaled counts (cm2)
UMG-3.3 unfolding package (GRAVEL & MAXED)
Neutron spectroscopy: unfolding
Allows for spectroscopy outside the laboratory
Requires accurate response functions … either simulated (< 20 MeV) or measured (>20 MeV)

## Slide 10: Fast neutron reactions on 12C
14 MeV
12C (n, n)12C
12C (n, n’)12C
12C (n, p)12B
12C (n, d)11B
12C (n, 𝛼)9Be
90 MeV

## Slide 11
National laboratory
Largest accelerator in Southern Hemisphere
K200 Separated Sector Cyclotron
25 - 200 MeV protons
6 - 26 MHz
https://tlabs.ac.za/accelerators/cyclotron-parameters/

## Slide 12
Fast neutron beam facility
at iThemba LABS

## Slide 13
Quasi-monoenergetic neutron beams at iThemba LABS
Neutrons from a 100 MeV proton beam on a 5 mm Li target
measured at 8.00 m
0o
16o
“difference” beam
Φ/Q ~2.5 x 104 cm-2 μC-1
Beam profile at 8 m at 0o
[~10 cm]
Full energy peak at (Ep – Q) is strongly forward biased
Low energy continuum from breakup reactions in target is approximately isotropic

## Slide 14
Pulse height L (MeVee)
Pulse shape S (arbitrary)
neutron-induced events
e
p
d
t
α
gamma ray-induced events
Time-of flight measurement of neutrons produced by a 66 MeV proton beam irradiating a 8.0 mm natLi target.
(BC501A at 8.00 m from the target at 0o).

## Slide 15
Measured light output spectra for neutron energies between 10 - 65 MeV, selected by time-of-flight for BC501A.

## Slide 16: Spectrum unfolding with measured detector response functions Kutullo Maibane (PhD)
Unfolded (MAXED)
Time of flight
Neutron light output spectrum measured at 0o with BC501A for a 66 MeV proton beam on 8.0 mm Li target

## Slide 17: Time of flight measurement of neutrons produced by a 66 MeV proton beam irradiating an 8.0 mm natLi target. (EJ-276 at 8.00 m from the target at 16o). Erin Jarvie (MSc)
Normalised counts
Energy (MeV)
BC-501A reference
EJ-276 compact
neutron-induced events
gamma ray-induced events

## Slide 18
Measured light output spectra for neutron energies between 10 - 65 MeV, selected by time-of-flight with EJ-276.

## Slide 19: Spectrum unfolding with measured detector response functions
Neutron light output spectrum measured at 16o with EJ-276 for a 66 MeV proton beam on 8.0 mm Li target

## Slide 20: Summary
Detector response functions were measured at the iThemba LABS fast neutron facility using time-of-flight between 10 - 65 MeV for the reference BC501A and EJ276 detector systems.
Unfolded energy spectra comparable to those measured by time-of-flight.
Compact spectrometer design update underway, but very promising for fast neutron spectroscopy outside of the laboratory!
Further measurements up to 200 MeV planned, with improved statistics to extend range of applications.

## Slide 21: Thank you!
This work was partially funded by the National Research foundation grant ID: 138069

## Slide 22: Pulse shape discrimination
Pulse shape parameter:
t(ns)
V(mV)
Short integral of pulse
Long integral of pulse
tL
tS
L: Light output, QL, scaled to MeVee using gamma ray sources
Pulse height L (MeVee)
Pulse shape S (arbitrary)
neutron-induced events
e
p
d
t
α
gamma ray-induced events

## Slide 23
Measured
[Goldhagen et al. Rad. Prot. Dos. (2004) vol.110 p.387]
Calculated
[Clem et al. Rad. Prot. Dos. (2004) vol.110 p.423]
Cosmic-ray neutron spectrum: 20 km at 54° N, 117° W
10.1140/epjp/s13360-021-01401-3