## Slide 1
The development of (70 ~ 100) MeV Mono-energetic neutron reference fields based on the 100 MeV Cyclotron
Jiaoting Yu, Wei Li, Shiyao Li, XI Qin, Xinqi Luo, Xueying Deng, An Du, Hailiang Qin, Yujun Mo, Yuyang He, Bin Shi, Shufeng Zhang, Yuntao Liu
Neutron metrology laboratory
Department of application of nuclear technology
China institute of atomic energy
12.9.2022, Beijing Cyclotron

## Slide 2
neutron metrology in CIAE
1
Why we need high energy neutron fields
2
C
O
N
T
E
N
T
CONTENT
Summary
5
Layout of high energy neutron field in CIAE
3
Experiment results
4

## Slide 3
Neutron Metrology and Measurement Laboratory
Neutron Metrology
Thermal neutron
Mono-energetic neutron (from 8 keV ~ 20 MeV)
Quasi-monoenergetic neutron
Now: 70 ~ 100 MeV
Future: 500MeV or 2GeV
White neutron field
Base on spallation neutron source
Neutron measurement in workplace
Cosmic ray measurement
Development of instruments for commerce
Particle beam Analyzing for materials
PIXE and PIGE
Positron annihilation
Neutron nuclear medicine
Neutron dosimetry for BNCT
Development of target for BNCT
micro /nano dosimetry
Neutron induced gamma ray analyze
Spectrum and dosimetry measurement on the workplace
Main fields of Neutron metrology Laboratory

## Slide 4
cold
thermal
Epi-thermal
Fast neutron
High energy
energy/MeV
fluence
0cm-2s-1
1015cm-2s-1
103cm-2s-1
106cm-2s-1
Thermal neutron from reactor
8keV~20MeV energetic neutron fields base on tandem accelerator
20
10-3
1
10-6
10-9
Graphite Thermal pile
100
500
100MeV quasi monoenergetic neutron fields base on cyclotron
500MeV quasi monoenergetic neutron fields base on linac
White neutron fields base on Spallation neutron source
Neutron reference fields
developing
developed
In the future

## Slide 5
neutron metrology in CIAE
1
Why we need QMN fields
2
C
O
N
T
E
N
T
CONTENT
Summary
5
Layout of QMN field in CIAE
3
Experiment results
4

## Slide 6
Why we need high energy neutron fields
Neutron will be produced by primary cosmic ray which is reacted with atmosphere, lunar soil, space station materials, astronaut…
For many purposes: such as astronauts healthy monitor, space environment monitor …;
all kinds of detectors were send to space: neutron personal dosimetry, TEPC, neutron spectrometer …;
Neutron energy distribution at different heights from the ground
Relative weights with neutron energy for dose equivalent
Tens of MeV neutron plays an important role!

## Slide 7
The key experiments for launching neutron detectors are: get fluence response, dosimetry response and energy calibration at interested neutron energy region.
Why we need high energy neutron fields
Launching neutron detector in China
< 20 MeV，MC simulation + monoenergetic calibration
> 20 MeV，MC simulation (only)
Status for now
problem
> 20 MeV，lack of cross-section data, depend on the theoretical model
above 20 MeV, measurement results quality is poor
result
Secondary neutrons in proton/hadron therapy
  Patient radiation protection (En max ~250 MeV / 800 MeV)  secondary cancer risk
Shielding benchmark experiments for high energy accelerator facilities
Cross section studies
  Accelerator-driven reactors, spallation sources
Study of neutron-induced single–event effect (SEE)
  Test of semiconductors

## Slide 8
Why we need high energy neutron fields
NMIJ (Japan)
  development of QMN fields at TIARA facility from JAEA
      calibration services at 45 MeV started in 2015
  Medium/long term : up to 400 MeV using RCNP facility
i-Themba Labs (South Africa)
  Development of QMN fields between up to 200 MeV
  Collaboration between i-Themba Labs, NMISA, PTB and IRSN
CIAE (China)
  QMN fields in development from 70 MeV to 100 MeV

## Slide 9
neutron metrology in CIAE
1
Why we need QMN fields
2
C
O
N
T
E
N
T
CONTENT
Summary
5
Layout of QMN in CIAE
3
Experiment results
4

## Slide 10
Layout of QMN in CIAE
General View of the 100 MeV Cyclotron
the current at ion source : more than 10 mA
Proton energy: （ 70 ~ 100 ）MeV

## Slide 11
Layout of QMN in CIAE
Proton extracted and transported in two common lines simultaneously (north and south lines)
6 beam lines
  North hall
    N1 beam dump/ beam line testing
    N2: radioactive ion-beam, inject into 13 MV tandem accelerator
    N3: produce isotope for medicine
  South hall
    S1: quasi monoenergetic neutron source
South hall
  S2: white light spectrum neutron source
  S3: proton radiation beam for radio-biological and single event effect testing of electronic device

## Slide 12
Layout of QMN in CIAE
Key points of QMN designing:
  fluence
    The higher the neutron fluence rate, the better；（related to the aperture of collimator, thickness of target）
  energy monochromaticity (sometimes has contradiction with fluence)
    The more monochromatic the energy, the better；（related to target materials, thickness, the aperture of collimator）
  Background proton（related to the bending magnet designing）
  Background neutron
    contamination neutron (proton react with magnet, beam pipe, beam stop foil ( Ti ), Ar gas in target chamber, air and etc.)
    Low energy neutron: 7Li(p,n)7Be Isotropic vaporized neutrons
    Room scattering neutron；
  Background photon
    Prompt gamma produced by proton reaction;（related to the aperture of collimator）
    Delayed gamma;（materials of collimator, shielding materials)
  Neutron and proton monitor（fission chamber, scintillation detector, faraday cup）

## Slide 13
Layout of QMN in CIAE
  Proton from two quadrupole magnets，faraday cup transports into target chamber, bending into beam dump (embedded in concrete shielding wall) by bending magnet.
  Collimators: two layers, inside: iron 1m ×1m×3m , 15 cm diameter aperture, outside concrete;
  Proton beam spot on target: 2 cm
  Distance from target to collimator: 2 m
  Distance from target to wall: 12 m

## Slide 14
Layout of QMN in CIAE
Proton energy /MeV | Lithium thickness/mm | Energy deposition on the target/MeV | Select thickness/mm
70 | 3 | 1.32 | 4
 | 4 | 1.76 | 
 | 5 | 2.21 | 
 | 6 | 2.66 | 
 | 7 | 3.11 | 
 | 8 | 3.56 | 
 | 9 | 4.02 | 
80 | 3 | 1.18 | 5
 | 4 | 1.58 | 
 | 5 | 1.98 | 
 | 6 | 2.38 | 
 | 7 | 2.79 | 
 | 8 | 3.19 | 
Proton energy /MeV | Lithium thickness/mm | Energy deposition on the target/MeV | Select thickness/mm
90 | 3 | 1.08 | 6
 | 4 | 1.44 | 
 | 5 | 1.81 | 
 | 6 | 2.18 | 
 | 7 | 2.55 | 
 | 8 | 2.92 | 
 | 9 | 3.29 | 
100 | 3 | 0.99 | 6
 | 4 | 1.33 | 
 | 5 | 1.66 | 
 | 6 | 2.00 | 
 | 7 | 2.34 | 
 | 8 | 2.68 | 
Simulate the thickness of target by SRIM
How to determine the thickness of the target

## Slide 15
Layout of QMN in CIAE
Target chamber
  Consists of Faraday cup, and target plate filled with Argon gas with 1.1 atm;
  five targets on the plate
    4, 5, 6 mm metal 7Li
    Fluorescent target
    Blank target

## Slide 16
  Simulation result (MCNPX, with CEM03.03 model which is combined the preequilibrium emission and evaporation model）
Layout of QMN in CIAE

## Slide 17
0 m
2 m
4 m
6 m
Layout of QMN in CIAE
Neutron Fluence cm-2s-1/uA | Neutron ambient dose equivalent mSv/h/uA | Peak energy /MeV | Peak neutron ratio | Photon fluence cm-2s-1/uA | photon ambient dose equivalent mSv/h/uA | Photon/neutron fluence
5.47E+04 | 7.19E+01 | 9.26E+01 | 50.29% | 4.79E+03 | 9.83E+01 | 8.77%
Simulation for 100MeV neutron beam profile
Simulation results for 100MeV neutron beam

## Slide 18
Proton energy/MeV | Range in graphite / mm | Range in Copper / mm
70 | 20.13 | 7.09
80 | 25.57 | 8.95
100 | 38.08 | 13.21
Layout of QMN in CIAE
Beam dump and bending magnet

## Slide 19
Faraday cup as proton beam monitor
Neutron monitor 1, U-8 fission chamber
Proton vacuum beam pipe
Neutron monitor 2, thin 5mm plastic scintillation detector
Layout of QMN in CIAE
Pulse height spectrum of plastic scintillator
Assembly of fission chamber assembly
The effective zone diameter of the coating is 25mm, and the mass thickness is about 200μg/cm2.

## Slide 20
neutron metrology in CIAE
1
Why we need QMN fields
2
C
O
N
T
E
N
T
CONTENT
Summary
5
Layout of QMN in CIAE
3
Experiment results
4

## Slide 21
Experiment results
238U fission chamber confirms that we get the neutrons---the first neutron experiment
238U FC experiment
Signals from the preamplifier by U-8 FC
Pulse height spectrum from U-8 FC
The main target material is natural U, the diameter of the target area is 100 mm, and the active aera is 78.54 cm2
238U FC

## Slide 22
Experiment results
Neutron monitors and proton monitor repeatability
  Fission Chamber around target: NM1
  Plastic scintillator in experiment room: NM2
  Faraday cup as the proton monitor: PM
  Neutron detector: N
Neutron energy/MeV | With/without target | Repeatability of N/NM1 | Repeatability of N/NM2 | Repeatability of N/PM
100 | With T | 1.86% | 1.48% | 3.45%
 | Without T | 1.67% | 1.01% | 1.09%
70 | With T | 1.04% | 0.77% | 1.41%
 | Without T | 6.67% | 5.11% | 4.02%
Repeatability of neutron and proton monitor

## Slide 23
Experiment results
Background
  neutron Background
    contamination neutron; ratio: N (with target) / N’ (without target), on beam
    Room scattering neutron; ratio: Ns (on beam) / N’s (out of beam), with target;
N / N’ \ distance (cm) | 23.1 | 53 | 100 | 300 | 500
70 | 3.27 | / | 2.01 | 1.82 | 2.29
100 | / | 2.63 | 2.60 | 2.48 | 2.27
Ns / N’s \ distance (cm) | 23.1 | 53 | 100 | 300 | 500
70 | 16.9 | / | 11.0 | 7.4 | 4.5
100 | / | 18.5 | 12.3 | 6.8 | 5.0

## Slide 24
Experiment results
Neutron beam profile
  Micromegas detectors, 512 channels, 64 + 64 strips (orthogonality) for readout , 10 cm × 10cm area per detector, 4 detectors for measurements.
  at 0.5 m, 1 m, 1.5 m, 2 m, 2.5 m, 3 m, 4 m positions.

## Slide 25
70 MeV
80 MeV
90 MeV
100 MeV
1 m
2 m
3 m
4 m
Experiment results

## Slide 26
Experiment results
gamma beam profile
7Li – 6Li pairs of thermoluminescence detectors
was placed 30 cm away from the outlet of the collimator
The XY plane was perpendicular to the direction of the beam.

## Slide 27
Experiment results
Neutron spectrum measurement
  DC beam mode, TOF method with two liquid scintillation detector, one for start, one for stop gate signal
  TOF detector1: EJ301, 2” diameter, detector2: EJ309, 5” diameter; coincidence electronic: CAEN 5790;
  Angles between secondary neutron and neutron beam line are 57°, 57°, 59° at 70 MeV, 90 MeV, 100 MeV
  Secondary neutron flight lengths are 3.46 m, 3.00 m, 3.00m at 70 MeV, 90 MeV, 100 MeV
  Detector2 use shielding B-PE to reduce the background of coincidence event.
Picture of two scintillation TOF experiment
Pulse height spectrum of detector1
Pulse height spectrum of detector2

## Slide 28
100MeV
90MeV
70MeV
TOF Spectrum
TOF Spectrum
TOF Spectrum
Neutron energy distribution
Experiment results

## Slide 29
Experiment results
Neutron fluence absolute measurement
  ΔE-E telescope, polyethylene convertor Φ38 mm×1 mm；ΔE detector: silicon, Φ40 mm×1.5 mm; E detector: BC501A Φ2’’ × 2’’; with two apertures to define the solid angle.
  Telescope is located outside the neutron beam to reduce the background events for ΔE-E detectors
  CAEN DT5790 is used for coincidence electronic box.

## Slide 30
Experiment results
Coincidence events TOF spectrum
Fluence response is calculated by Geant4
Energy/ MeV | 70 | 80 | 90 | 100
Response/ cm2 | 1.03E-05 | 1.51E-05 | 2.02E-05 | 2.92E-05
Fluence /cm-2 s-1μA-1 | 8.55E+04 | 7.12E+04 | 6.42E+04 | 5.04E+04
Coincidence events pulse height spectrum for silicon
Coincidence events pulse height spectrum for silicon
Measurement point: 80 cm from outside of collimator
compare to the simulation result, 24% relative error

## Slide 31
neutron metrology in CIAE
1
Why we need QMN fields
2
C
O
N
T
E
N
T
CONTENT
Summary
5
Layout of QMN in CIAE
3
Experiment results
4

## Slide 32
Summary
QMN is important for detector calibration, SEE, proton therapy, shielding materials benchmark, cross-section measurement;
70 MeV ~ 100 MeV QMN is established in CIAE for the first time in China
Get the information of QMN
  Neutron monitors: FC, thin plastic scintillation;
  Proton monitor: faraday cup;
  Neutron beam profile: micromegas detector;
  Energy: two scintillation TOF method;
  Fluence: ΔE-E telescope
Next plan: reduce the neutron background
  Reduce the contamination neutrons
  Fully close the neutron target Chamber by concrete shielding materials
  Add the neutron beam dump to reduce the scattering neutron.

## Slide 33
Thanks for your attention